FEI TECNAI T-12 SPIRIT AND TWIN

GENERAL USE INSTRUCTIONS

THIS MANUAL IS TO BE USED AS A GUIDELINE FOR GENERAL USE OF THE TECNAI T-12. YOU ARE EXPECTED TO TAKE NOTES AND KEEP A NOTE BOOK AS YOU PROGRESS THROUGH USING THE S/TEM AS A RESEARCH TOOL. PLEASE BRING THESE INSTRUCTIONS AND YOUR NOTEBOOK WITH YOU EACH TIME YOU USE EITHER T-12.

RULES FOR USING DUFFIELD 150 AND SB56 BARD

WE NEED TO SET STRICT GUIDELINES FOR USING THIS FACILITY SO THAT WE CAN MAINTAIN A CLEAN AND ORGANIZED RESEARCH ENVIRONMENT. SO THAT YOU GET INTO GOOD S/TEM HABITS WE WILL BE TREATING ALL SAMPLES UNIFORMLY AND ALL USERS WILL ADHERE TO THESE LAB RULES. IF YOU HAVE NOT DONE SO ALREADY, YOU MUST HAVE YOUR ID CARD PROGRAMMED SO THAT YOU CAN HAVE ACCESS TO EITHER LAB (150 DUFFIELD AND SB56 BARD). YOU NEED TO GO TO THIS WEBSITE: http://www.ccmr.cornell.edu/facilities/safetyinfo/, AND CLICK ON THE APPROPRIATE FACILITY YOU NEED ACCESS TO. THE CLASSES THAT YOU NEED FOR ACCESS ARE LISTED, FILL OUT THE FORM AND SEND THE FORM TO jlg98 OR djw326.

- NO FOOD OR DRINK IN EITHER LAB

- YOUR SHOES MUST BE CLEAN UPON ENTERING THE MICROSCOPE ROOMS. BRING A CLEAN PAIR (NO OPEN TOE SHOES).

- NEVER ATTEMPT TO REPAIR ANYTHING IN THE FACILITY, CONTACT JG IMMEDIATELY IF THERE IS A PROBLEM. SOFTWARE AND COMPUTERS ARE INCLUDED.

- NEVER USE YOUR OWN MEDIA (MEMORY STICKS INCLUDED) TO ARCHIVE IMAGES, NEW BLANK CDs ARE PROVIDED, YOU MAY USE DROP BOX ON THE SPIRIT OR THE BIOTWIN

- BRING YOUR NOTEBOOK AND INSTRUCTIONS EACH TIME YOU USE THE MICROSCOPES

- SINCE THE TOOLS PRODUCE X RADIATION YOU MUST READ THE EHS X-RAY MANUAL AND SIGN THE WAIVER BEFORE YOU USE THE MICROSCOPES

- CHECK THE WEB SCHEDULE BEFORE YOU USE ANY OF THE MICROSCOPES USING THE CORAL SYSTEM

- TAKE YOUR SAMPLES WITH YOU WHEN YOU ARE DONE AND LEAVE THE LAB IN A CLEAN STATE

PROPER SAMPLE HANDLING PROTOCOL

Don’t worry, you can look at pretty much anything in these microscopes, but the samples must be extraordinarily clean before they go into the microscopes. This protocol is not only for the welfare of the microscope, but if your sample contaminates you will have wasted your time. Make sure that when you prepare your samples that your tweezers are cleaned with a solvent such as methanol, your grids have been cleaned with a solvent and you must have a clean grid box (you can clean these with ethanol). If you wish, use gloves when you prepare or handle your samples, if not just don’t touch the business end of the tweezers or the grids with your fingers

.

If you are looking at inorganic samples in the SPIRIT the sample and the sample holder must be cleaned in the plasma cleaner for three minutes.

If you are looking at organic samples that have a good track record* under the beam you just need to clean the single-tilt holder for three minutes before you load the sample. If your sample historically contaminates you must clean your samples in the plasma cleaner for 5 seconds. If that does not work place your sample under UV for thirty minutes before you load the sample in the plasma cleaned single-tilt holder.

Any sample that hits the floor must be cleaned in either the plasma cleaner, under the UV light or just ditch it.

We do not have to plasma clean the holder for the TWIN in Bard, BUT please do not touch the holder, if you do let the manager know and we will show you how to properly clean the holder

The cold trap in each microscope must be filled before you load your samples. Once filled, it will take at least twenty minutes before you can load a sample. This is usually done by me first thing in the morning; the dewar must be checked every four hours though. If you are the last user of the day you must put the microscope through a cryo-cycle before you leave.

*Has not contaminated in our other microscopes

(1)

GETTING STARTED WITH THE T-12s

After logging onto Coral, log onto your account which will be assigned to you during the training session; you will have one account on the service computer and another on the support computer. When you come into the scope room the RED button on the console will be illuminated indicating that the T-12 is on. The monitors will be at the log on prompt. Type in your user name and password. If there are any issues with the scope during the session please note them in Coral and send the manager an email

- When the main page opens up, go to the menu bar at the bottom of the monitor and open the programs in this order: Tecnai user interface, and analySIS (if you are doing STEM or x-ray analysis you will need to open up TIA, if not leave it closed). In order for the cameras to work properly you must open the programs in this order.

- Under WORKSET you will be at SETUP.  Take note of the Gun/Col vacuum value, the log value should be 6. Note that the Col. Valves Closed, Filament and High Tension buttons are all YELLOW, this indicates that the TEM is operational

- Next to TECNAI G2 on the lower Right of the first monitor open up Vacuum Overview in the pull down. This diagram shows you a cut away of the vacuum system and all the values.

- If the dewar is not on the scope, fill the dewar with liquid nitrogen and place the copper braids in the dewar. Place the Styrofoam lid on top of the dewar and wait twenty minutes. If the dewar is on you can get started right away.

- Look in the lower Right corner of the monitor, you will see X, Y, Z and A, B. They should all be at ZERO. If not go to SEARCH in WORKSET. There will be an arrow at the top of STAGE, click it to get the flap out. You can now go to RESET and click HOLDER; this will zero the goniometer.

- Return to SETUP.

- You are now ready to load your sample.

(2)

LOADING YOUR SAMPLE

This should be a pretty simple protocol, but over the years I have found that most of my scope problems occur when this procedure is done incorrectly. You have to do this properly each time, so take your time both loading and unloading your sample; think about what you are doing and if anything is out of the ordinary call me immediately.

- Locate the T-12 holders and choose a holder. We have a single and a double tilt holder for the Spirit, and a Single and a Cryo holder for the Twin. We usually have the single tilt holder in the column, so it has to be removed first.

- If you are looking at organic material, or samples that does not need the Alpha/Beta tilt you must use the single tilt holder.

- FOR THE SPIRIT TWIN ONLY: Cleaning the Holder: Make sure that the valve on the Argon/Oxygen tank is open and then switch the Plasma Cleaner ON (switch on back panel next to plug).  The Plasma Cleaner will automatically vent to atmosphere, this takes about three minutes. Once at atmosphere, remove the white vacuum plug and carefully replace the plug with the sample holder of choice. Now press the vacuum button. While the chamber is evacuating you can set the time for three minutes and then you hit the set button; if you don’t hit the set button nothing will happen. The cleaner will activate the plasma cloud once high vacuum is reached. After cleaning press the VENT button and wait until the cleaner is at atmosphere. Remove the now cleaned holder and put it in its cradle. Replace the sample holder with the white vacuum plug and evacuate the chamber. When the chamber is at high vacuum you may turn the unit off.

- Using the Clamp opening tool, place the end into the hole of the sample clamp. Gently lever the clamp up to a 90 degree angle and remove the tool. Place your grid in the depression, sample side down, with clean tweezers, re-insert the clamp tool and lever the clamp onto your grid.

- Inspect the o-ring for debris. If you find debris, gently sweep the debris away with the tweezers, never grab at the junk, this can damage the o-ring.

- Make sure that the vacuum overview window is up, this will show you what the pumping system is doing and when the pump sequence is over.

- READ THIS SECTION BEFORE YOU DO ANOTHER THING! Load the sample in the TEM. Align the pin on the holder to the line at the five o’clock position on the goniometer. You have now reached the point of no return; insert the holder carefully until it makes contact with the flange. You may or may not have hit the slot you need to be in. Carefully wiggle the holder no more then 1 or 2 degrees to feel for the slot. The PVP pumps on the goniometer for 90 seconds.  The PVP will pump on the goniometer for 90 seconds. In the highlight area the microscope will ask you which holder you have selected, you must tell the microscope which one (single or double-tilt) or you will have no control over sample movement. If you choose the wrong holder you will have to remove the holder and reinsert so that the scope knows which holder you are using. When the PVP has stopped pumping you will carefully rotate the holder counter clockwise until it stops; the interlock is now open to the high vacuum. NEVER LET THE HOLDER GO! GUIDE THE HOLDER INTO THE COLUMN CAREFULLY UNTIL IT STOPS.

- You must wait until the vacuum in the GUN/COL returns to a value of less than 10 before you open the COLUMN VALVE. If you don’t use LN2 you will not reach the ultimate vacuum of 6 for a long-long time.

ALIGNMENT OF THE TEM

The alignment of the T-12 is controlled by the computer. Because of that the best way to align the microscope is to use the alignments that I do once a week for 120, 80 and 60kV.

- Make sure that the scope is at 2400x on the SPIRIT and 1050x on the TWIN, both need to be in TEM Bright Field

- Open the column valve. Press the Yellow Col. Valves Closed button, you will hear the valve open and you should see the beam. If not, move your sample around a bit using the track ball on the RHS. IF YOU NEED TO LEAVE THE ROOM FOR ANY REASON YOU MUST MAKE SURE THAT THE COL. VALVE IS CLOSED.

- Find something on your sample that is easy to distinguish. Even though we hammer into you guys that your samples need to be extremely clean, we need schmutz (a piece of junk) to set the eucentric height.

- Press the eucentric height button on the RHS control pad. This sets the focus voltage to the objective lens and should be near “focus”.

- Once you have found your schmutz, go to SEARCH in the WORKSET. Open the flap out under STAGE2 and activate WOBBLER. You will notice that your schmutz is moving. Using the Z AXIS controls on the RHS control panel move the sample either up or down until the movement has minimized. Increase your magnification to around 20,000x to fine tune your adjustment.

- You can now deactivate the WOBBLER.

- Lower the FOCUS SCREEN and Focus your image. This is very important because all your alignments are going to be based on your focus, your focus % will be around 93 for the SPIRIT and 90 for the TWIN.

- Open the Tune or the Align tab. In the upper right corner expand the window and select FILES. Look for the newest alignments with the appropriate kV. Click on the file and you will see a set of AVAILABLE alignments. Select all of the alignments and move them to the SELECTED. Then hit apply and the microscope is aligned.

HOW TO TAKE AND ARCHIVE AN IMAGE IN THE SPIRIT

- Look around your sample until you are satisfied that you have sample and that it is worthy of analysis. Try not to waste your time on a useless sample.

- You probably need contrast at this point. If you don’t just go to the next step. Find material that is not valuable and press the DIFFRACTION button on the RHS. Insert the OBJECTIVE APERTURE ASSEMBLY by flipping the silver lever to the left. You should see the image of the OBJECTIVE APERTURE around the DIFFRACTION spot. There are 4 objective apertures, 10, 20, 40 and 100 micron. Choose the one that best suits your sample by turning the large knob on the assembly. Using the smaller knobs you can move the aperture X and Y making the aperture concentric around the diffraction spot. Again, these are NEVER TOO FAR OFF. If you find yourself out of control and sweating profusely call me immediately.

- Depress the DIFFRACTION button to return to BRIGHTFEILD

- Using the FOCUSING SCREEN, focus your image using the binoculars

- Spread the beam until it fills the large screen. YOU MUST MAKE SURE THE INTENSITY KNOB IS AT ITS FINEST SETTING.

- Insert the camera. Next to the control pad on the RHS you will see a silver mouse like thing with a red and green light on it. When you press the green light you will hear a thunk, that is the camera being inserted into the path of the beam, the red light will remove the camera. Press the green button.

- You will now be using the monitor on the RHS which is the SIS software. With the camera inserted you can now activate the camera by pressing the camera icon with the #3 in it on the upper menu bar. When you do this, a histogram will pop up, a live FFT and the command for the FFT.  There should be an image on the screen at this point and the histogram will be active. CAREFULLY adjust the histogram with the intensity knob; you should be near 50% FOR A 1SECOND EXPOSURE. If your image looks good press the INTENSITY ZOOM button, the intensity will be locked through your magnifications. If you find you need more intensity, increase it, but you must unlock and then lock the intensity again. If at any time you are looking at really contrasty images you may need to bail on the auto-histo setting and do it manually, it is very good practice none-the-less.

- Under IMAGE, scroll down to CAMERA CONTROL, double click, you now have control of the camera speeds. Choose 200ms for cruising around your sample.  This is where you can focus and stigmate your image using the FFT, please give it a try, it really works and makes your images look much better.

- ACQUIRE your image. Once your are focused and stigmated, set your camera speed to at least 1 second, press the icon with the #3 in it and your image is now archived. Press the UNION JACK ICON and the scale bar will be burned into the overlay.

- SAVE your image. You will have an account on the support PC, which is the “Z” drive. You must use this PC and, of course never use a memory stick.

- We use XP burn to save your images to CD. You can also use DROP BOX OR GMAIL. Please do not use this PC as a mass storage device. I do delete images…eventually.

HOW TO TAKE AND ARCHIVE AN IMAGE IN THE TWIN

KEEP IN MIND THAT THE TWO CAMERA SYSTEMS ARE VERY DIFFERENT. DIGITAL MICROGRAPH (DM) IS THE INDUSTRY STANDARD AND YOU WILL SEE IT AGAIN. DM IS A VERY DEEP PROGRAM AND IS NOT ALL THAT USER FRIENDLY, BUT AGAIN, EVERYONE “OUT THERE” USES IT AND IT DOES DO A FANTASTIC JOB.

- Make sure the camera is cooled down: Behind the Monitors is the Camera Box. It has six lights on it. Four of the six will be green, the fourth will be off and the sixth will be orange; that indicates that the camera is out (the fourth light is camera insert) and the camera is cooling down (the sixth light orange). Once the sixth light is green the camera is cooled down.

- Now that DM is ready, in the upper RHS of the support screen you will see the camera controls. Under camera view you can insert the camera by clicking the box.

- You now select search under setup

- Next select start view and the image will appear on the screen

- On the LHS of the screen you will see the histogram window. Move your curser onto the image for the histogram and the counts. The number that appears at the bottom of the histogram window should not exceed 6000e- or you will damage the camera. If you are in that range lock the intensity limit.

- To focus, select focus under camera view.

- To stigmate the objective lens using the FFT, select the dashed box on the upper menu bar. Hold the ALT key down and draw a box; the box must be red. This will create the necessary 2k x 2k box. Under process select FFT and then Live FFT. The FFT will then appear in the upper RHS of the screen. To correct the objective stig, select an amorphous area, adjust the focus to expand amorphous pattern and then using the x,y-multifunction controls adjust the pattern so it is circular.

- To save an image, click image acquire in the camera acquire box. Check your image and then go to file and save as. Select the format you want and save.

- Burn your CD using XP burn;  or DROP BOX OR GMAIL.

SHUTTING DOWN

MAKE SURE THAT YOU SHUT THE SCOPE DOWN IN THIS ORDER EACH TIME.

- Make sure that your images are saved and burned to CD.

- RETRACT THE CCD CAMERA FROM THE COLUMN

- Lower the magnification to 2400x for the SPIRIT and 2550x for the TWIN, the scope needs to be in TEM BRIGHTFEILD

- Spread the beam so that it fills out the screen

- Remove the OBJECTIVE APERTURE

- Go to SEARCH in WORKSET and RESET the holder

- Close the COLUMN VALVES

- Bring up VACUUM OVERVIEW

- Remove the sample holder. With your Left hand press against the purple faceplate while you pull the holder out of the column. When the holder can be pulled no further rotate the holder CLOCKWISE until it stops. You should now gently “POP” the holder out of the goniometer, and remove the holder to the cradle in the other room.

- Record your time in Coral

- Log off the computer. Save your settings.

- YOU MUST LEAVE THE ROOM IN THE ORDER WHICH YOU FOUND IT, THERE IS NO MAID SERVICE.

STARTING UP AND SHUTTING DOWN FOR THE NIGHT

Remember that you are not allowed to use the scope after hours or on the weekend until you have had 20 hours of uneventful microscopy.

START UP

- Log on to the PC and start UI

- Put Liquid Nitrogen on the ACD (anti contamination device); wait at least 15 minutes before loading your sample

- Make sure the vacuum values are good

- Click the HT tab, you need to hear a high pitched whistle sound, if not click the HT off and then back on again

- Click the Filament tab and wait approximately 300 seconds for the filament to heat up

- Load your sample and start breaking down the barriers of science

SHUT DOWN

- When your last sample is removed from the scope and the scope is at    the defaults you may now start shutting the scope down for the night

- Make sure there are no other users on after you (check Coral)

- Open the flap out in VACUUM under SET UP

- Find the CRYO CYCLE tab and click it

- Place the cryo mat under dewar and then remove the dewar from the copper braids.

- The scope will be in cryo cycle for 240 minutes so make sure that it will be ready for the next user.

- Go home